A New Approach for Regulating Oxygen Levels in a Microfluidic Device

Peter C. Thomas, a,b Srinivasa R. Raghavanb,c and Samuel P. Forry*a

Received (in XXX, XXX) Xth XXXXXXXXX 200X, Accepted Xth XXXXXXXXX 200X

First published on the web Xth XXXXXXXXX 200X

DOI: 10.1039/b000000x

Oxygen levels play a critical role in cell culture and have been shown to impact cell growth, death and differentiation. As such, proper management of the oxygen level is required in cell based assays. Poly(dimethylsiloxane) (PDMS) based microfluidic devices are gas permeable and offer the advantage of fast and accurate oxygen regulation compared to conventional incubators. However, maintaining proper gas levels requires continuous gas flow which leads to pervaporation of water through the PDMS and, in turn, alters the osmolarity in the cell culture medium. In the present study, we demonstrate an alternative method for on-chip oxygen control. Instead of gases, we use aqueous solutions with different dissolved oxygen levels to regulate the oxygen content in the microfluidic test chamber. An off-chip gas exchanger is used to equilibrate each control solution prior to entering the chip. Using this strategy, problems due to pervaporation are considerably reduced. An integrated PDMS-based oxygen sensor allows accurate real-time measurements of the oxygen within the microfluidic chamber. The measurements were found to be consistent with predictions from finite-element modelling.

Introduction

Oxygen is a critical parameter for cell survival and homeostasis. While oxygen is required for ATP production, too much or too little oxygen greatly changes cell response and behavior. In vitro studies have demonstrated that oxygen levels affect cellular growth,1 enzyme expression2 and stem cell differentiation.3, 4 Also, oxygen gradients as a result of mass-transfer limitations occur naturally in normal and developing tissues as well as within growing tumour masses. These oxygen gradients have been suggested to regulate proper tissue function,5 influence cell development,4 and elicit the onset of angiogensis6. To improve our understanding of oxygen-dependent cellular behavior, proper control of oxygen level during cell culture is critical.

In general, mammalian cells in tissue are found at an oxygen partial pressure (PO2) of 3% to 9% oxygen (PO2 = 0.03 atm – 0.09 atm).7 In contrast, in conventional cell culture, cells are maintained at the ambient level of 21% oxygen (PO2 = 0.21 atm). To reduce the oxygen level during in vitro cell culture, various hypoxic chambers have been developed. However, the time required for culture medium to reach equilibrium and the lack of ability to generate oxygen gradients limit the scope of application for such chambers.8

Gas-permeable, poly(dimethylsiloxane) (PDMS)-based microfluidic devices offer alternative approaches to existing methods of oxygen regulation.9-14 Utilizing passive diffusion, external oxygen gas pumped through a control channel equilibrates through PDMS with the dissolved oxygen in adjacent fluidic channels, thereby achieving oxygen regulation on the device. Given the small distance between channels and the high permeability of oxygen through PDMS15, equilibrium can be achieved quickly (within seconds).12, 13 In addition, oxygen gradients can be generated using different oxygen-nitrogen gas mixtures across a device.13 While this method of oxygen control is an improvement, the current design allows significant pervaporation of water through the PDMS.16 In mammalian cultures where the temperature is held at 37°C, the aqueous liquid within a fluidic chamber can quickly evaporate and pass through the PDMS, resulting in dramatically increased osmolarity of the liquid.16 This issue is potentially exacerbated when gas-filled control channels are placed in close proximity to fluidic channels for gas control. To circumvent this problem, mammalian cells are often cultured in the device under continuous flow conditions. However, this precludes culturing cells that are sensitive to shear stress and limits the potential applicability of microfluidic methods.

In the current study, we demonstrate a new approach to regulating oxygen on a microfluidic chip. In contrast to previous methods, oxygen levels in stagnant on-chip chambers here were modulated by moving aqueous solutions with different dissolved oxygen levels through neighboring control channels. Utilizing an off-chip gas exchanger, the control solutions are first equilibrated with the gas in the gas exchanger and subsequently flowed through the device. We demonstrate that this method can effectively regulate the oxygen level and also significantly reduces the effect of pervaporation. In addition, we use an oxygen sensor integrated into the PDMS device17 to monitor the changing oxygen levels within the fluidic chamber continuously and in real-time.

Experimental Methods

PDMS-based multilayer microfluidic devices were fabricated using photolithography and soft lithography techniques.18 The current design has a single fluidic line that flows into a cell culture chamber, which is positioned between two control lines (Figure 1). Pneumatic valves were placed upstream and downstream from the culture chamber to regulate the liquid flow. Control lines and the culture chamber were 250 µm and 1000 µm wide, respectively. Two designs were tested: in one, the horizontal distance between the culture chamber and the control line was 200 µm and in the other this distance was 80 µm. The vertical distance between the two was ~ 10 µm in all cases. All channels were 30 µm tall. The gas exchanger was constructed using gas-permeable Teflon AF tubing with i.d. = 600 µm and o.d. = 800 µm (Biogeneral Inc, San Diego, CA).

Figure 1 Schematic of microfluidic device setup for on-chip oxygen regulation. Water was pumped through the gas exchanger where the PO2 in the liquid is reduced via equilibration with the controlled gas composition. This solution then enters the microfluidic control lines flanking cell chamber of the device. An oxygen sensor integrated into the floor of the device (red) measures changes of PO2 as dissolved oxygen in the control lines diffuse through the PDMS and into the adjacent cell chamber. Schematic not drawn to scale.

An oxygen-sensing PDMS film, based on the porphyrin dye, Pt(II) meso-tetrakis(pentafluorophenyl) porphine (PtTFPP, Frontier Scientific Inc, Logan, Utah), was prepared and calibrated as described previously.17 Briefly, PtTFPP was dissolved in PDMS (Slygard 184; Dow-Corning, Midland, MI) and spin coated onto a microscope slide. Separate layers of Teflon AF (Dow Chemical) and PDMS were sequentially spin-coated and cured. The total sensor thickness was ~ 24 µm. This oxygen sensor was integrated with the microfluidic device through plasma bonding. Calibration of the on-chip sensor was accomplished by exposing the device to different oxygen levels. Emission intensity from the sensor was captured and used to generate a calibration curve. Because the thin-film sensor covered the whole device, spatial variations in oxygen level could be monitored across the entire device.

Pervaporation experiments were performed at a temperature of 37°C. A 10 µM solution of dextran conjugated with rhodamine (molecular weight ~70,000; Sigma, St. Louis, MO) was used as a marker for changes in the solute concentration. A steady-state 2D finite element model was developed using FlexPDE software based on the device and channel geometry (Figure S-2).

Results and Discussion

Regulation of oxygen level on-chip was accomplished using an off-chip gas exchanger. Water, pumped through the gas exchanger, first equilibrated with the oxygen level there. The high oxygen permeability of Teflon tubing (990 barrer, from manufacturer website) and the short diffusion distance (tubing wall thickness ~ 100 µm) allowed rapid equilibration with the PO2 set point. At a flow rate of 700 µL/min (residence time of the liquid in the gas exchanger was ~100 s), the oxygen level in the liquid had fully equilibrated by the time the liquid left the gas exchanger.

Previously, we had developed a PDMS-based, thin-film oxygen sensor using the phosphorescent dye, PtTFPP.17 Here, we integrated this sensor film into the floor of the microfluidic device through plasma bonding. The emission intensity from the film is quenched in the presence of oxygen, and the relationship between the two is given by the linear Stern-Volmer equation:19

IoI=1+KSV×PO2	(1)

where I0 is the phosphorescence intensity of the sensor under pure N2 (PO2 = 0 atm), I is the quenched intensity at elevated oxygen levels, KSV is the Stern-Volmer constant and PO2 is the partial pressure of oxygen. The calibration of the sensor was performed via pixel-by-pixel analysis where each pixel was assigned a separate Stern-Volmer constant. The KSV determined for the on-chip sensor (517 atm ± 100 atm) matched previous bulk sensors17, indicating that the sensitivity of the measurement was not affected by the assembled microfluidic device (Figure S-1). As liquid carrying lower dissolved oxygen was introduced on-chip, the oxygen level within the device equilibrated by diffusion and was detected by the oxygen sensor.

Given the small volume of liquid within the microfluidic device, pervaporation can rapidly alter solute concentrations, potentially affecting a cell culture experiment. To monitor the changes in solute concentration during operation, emission intensity of dextran conjugated with rhodamine was monitored continuously in the fluidic channel while a downstream valve was closed. When warm humidified air was pumped through the control lines, an increase in rhodamine fluorescence with time was observed in the stagnant fluidic channel (Figure 2a). This was consistent with fresh rhodamine solution entering the chamber to replace liquid lost through pervaporation. From the steady increase in solute concentration, an osmolality doubling time of 2 h was calculated.  For the geometry of the system, this gave a pervaporation rate of ~0.25 nL/min (ESI). In contrast, pervaporation effects were minimized when warm liquid was pumped through the control lines, and no increases in rhodamine intensity were observed (Figure 2b). The use of water-filled control lines here is analogous to the incorporation of a water reservoir on the device to prevent pervaporation.20

By reducing the PO2 in the control lines, oxygen levels can be reduced across the fluid channel. The thin film sensor integrated into the floor of the microfluidic system allowed spatial variations in PO2 across the entire device to be monitored (Figure 3a).  Previous devices demonstrated oxygen regulation by directly overlapping the gas and fluid channels.12, 13 Adler et al. demonstrated the ability to generate different oxygen gradients by designing multiple control lines directly above a liquid line.13 Equilibration times in these systems varied depending on the distance between channels and were typically on the order of seconds. The current device minimizes potential light scattering and optical interference by incorporating control lines adjacent to the fluidic chamber.

Figure 2 Effects of water pervaporation during device operation. Fluorescent intensity of dextran conjugated rhodamine solution increased due to pervaporation when humidified air was pumped through the control lines (a). On-chip regulation of gas partial pressure using water-filled control lines (white circle) mitigated the effects of pervaporation as compared to humidified gas flow (black circle) (b). A pervaporation rate of 0.248 nL/min was determined from calculation based on using humidified gas flow and the device geometry. Pervaporation experiments were performed at 37oC. Data points represent normalized mean intensity value across the cell chamber (triplicates) with standard deviation indicated by error bars. Solid lines show the linear regression for each system with the fitted equations.

With the current system, steady state in the control lines was achieved in 2 min after air was switched to N2 in the gas exchanger (Figure 3b). The oxygen level in the cell chamber equilibrated more slowly and reached a steady state in 8 min. A PO2 across the whole device of 0.03 to 0.05 atm was achieved with this design (Figure 4a), providing a physiologically relevant oxygen enviornment.  A 2D finite element model was developed based on the steady state diffusion equation:

DO2∂2C∂x2=0			    (2)

Figure 3 Rapid PO2 response on-chip.  Phase contrast (left) and PO2 (right) images of the microfluidic device show the position of cell chamber with two flanking control lines (a). Initially, water equilibrated with air (from the gas exchanger) flows through the control lines, and the PO2 in the stagnant fluid chamber remains at 0.2 atm.  When air is switched to N2 (PO2 = 0 atm; t = 0 min), oxygen levels in the control line reached steady state within 2 min, matching closely the residence time of liquid in the gas exchanger (b). Steady state oxygen level in the cell chamber is achieved in 8.5 minutes.  The PO2 image is taken from t = 20 min in (b).

where DO2 is the diffusion coefficient of oxygen in PDMS, c is the oxygen partial pressure and x is the distance.  Based on the cross section of the device, simulations were performed to determine the on-chip oxygen levels as solutions with different dissolved oxygen content were pumped through the control lines (Figure S-2).  The height of the device was the only adjustable parameter for the model.  The simulation showed good agreement with the experimental data (Figure 4a) indicating oxygen transport in the device can be explained with a simple 2D diffusion model from eqn (2).

By pumping air-equilibrated and N2-equilibrated water into opposing control lines, an oxygen gradient was formed across the device (Figure 4b). The finite element model again showed good agreement with experimental data.

Conclusions

Precise regulation of oxygen level during cell culture is critical to understanding how oxygen affects cell function and response. Gas permeable PDMS based microfluidic devices offers the advantage of fast and precise control of oxygen level. However, continuous operation can result in pervaporation and dramatic changes in solute concentration that negatively affect cell behavior. We have developed a microfluidic cell culture system that accurately measures and controls oxygen levels while mitigating pervaporation. An oxygen sensor, integrated into the floor of the microfluidic device allows real time monitoring of the oxygen level on-chip. Instead of pumping gas through the device, liquids with varying oxygen concentrations, controlled with an off-chip gas exchanger, were pushed through the control lines to achieve regulation of oxygen level. We demonstrated that physiological oxygen levels (PO2 of 0.03 atm to 0.05 atm) can be maintained on-chip. Oxygen gradients were also generated. Experimental PO2 measurements showed excellent agreement with a finite element simulation. The current strategy allows the oxygen level to be monitored across the entire device, thus providing a tool for understanding how changes in oxygen level propagates throughout the entire microfluidic system.

Figure 4 PO2 equilibration. The oxygen level across the device equilibrated with various PO2 set points in the solutions pumped through the control lines (a). By flowing air equilibrated and N2 equilibrated water on opposite sides of the cell chamber, an oxygen gradient was formed across the device (b). Oxygen measurements from the sensor match the result of a finite element simulation (black line). Grey regions in the graphs indicate the location of the two control lines flanking the middle cell chamber, separated by PDMS membrane in white.  Data points represent the mean PO2 value with standard error shown with error bars.

a Biochemical Science Division, National Institute of Standards and Technology, Gaithersburg, MD, 20899,USA. Tel: 301-975-5246; E-mail: samuel.forry@nist.gov

b Fischell Department of Bioengineering, University of Maryland, College Park, MD,20742,USA .

c Department of Chemical and Biomolecular Engineering, University of Maryland, College Park, MD,20742-2111,USA

†	Electronic Supplementary Information (ESI) available: calibration of on-chip oxygen sensor (Fig. S-1), Pervaporation calculation (ESI) and numerical simulation of oxygen level in microfluidic device (Fig. S-2). See DOI: 10.1039/b000000x/

‡	Certain commercial products are identified in order to adequately specify the experimental procedure; this does not imply endorsement or recommendation by NIST.

Notes and references

1.	S. Parrinello, E. Samper, A. Krtolica, J. Goldstein, S. Melov and J. Campisi, Nat. Cell Biol., 2003, 5, 741-747.

2.	J. P. Camp and A. T. Capitano, Biotechnol. Prog., 2007, 23, 1485-1491.

3.	B. Keith and M. C. Simon, Cell, 2007, 129, 465-472.

4.	M. C. Simon and B. Keith, Nat. Rev. Mol. Cell Biol., 2008, 9, 285-296.

5.	K. Jungermann and T. Kietzmann, Hepatology, 2000, 31, 255-260.

6.	P. Carmeliet and R. K. Jain, Nature, 2000, 407, 249-257.

7.	B. Sahaf, K. Atkuri, K. Heydari, M. Malipatlolla, J. Rappaport, E. Regulier, L. A. Herzenberg and L. A. Herzenberg, Proc. Natl. Acad. Sci. U. S. A., 2008, 105, 5111-5116.

8.	W. E. Wright and J. W. Shay, Nat. Protoc., 2006, 1, 2088-2090.

9.	A. P. Vollmer, R. F. Probstein, R. Gilbert and T. Thorsen, Lab Chip, 2005, 5, 1059-1066.

10.	J. Park, T. Bansal, M. Pinelis and M. M. Maharbiz, Lab Chip, 2006, 6, 611-622.

11.	R. H. W. Lam, M. C. Kim and T. Thorsen, Anal. Chem., 2009, 81, 5918-5924.

12.	M. Polinkovsky, E. Gutierrez, A. Levchenko and A. Groisman, Lab Chip, 2009, 9, 1073-1084.

13.	M. Adler, M. Polinkovsky, E. Gutierrez and A. Groisman, Lab Chip, 2010, 10, 388-391.

14.	J. M. Higgins, D. T. Eddington, S. N. Bhatia and L. Mahadevan, Proc. Natl. Acad. Sci. U. S. A., 2007, 104, 20496-20500.

15.	T. C. Merkel, V. I. Bondar, K. Nagai, B. D. Freeman and I. Pinnau, J. Polym. Sci. Pt. B-Polym. Phys., 2000, 38, 415-434.

16.	Y. S. Heo, L. M. Cabrera, J. W. Song, N. Futai, Y. C. Tung, G. D. Smith and S. Takayama, Anal. Chem., 2007, 79, 1126-1134.

17.	P. C. Thomas, M. Halter, A. Tona, S. R. Raghavan, A. L. Plant and S. P. Forry, Anal. Chem., 2009, 81, 9239-9246.

18.	M. A. Unger, H. P. Chou, T. Thorsen, A. Scherer and S. R. Quake, Science, 2000, 288, 113-116.

19.	J. R. Lakowicz, Principles of Fluorescence Sepctroscopy, Plenum Press, New York, 1983.

20.	J. W. Song, W. Gu, N. Futai, K. A. Warner, J. E. Nor and S. Takayama, Anal. Chem., 2005, 77, 3993-3999.