ARE BRAGG PEAKS GAUSSIAN?

Boualem Hammouda

Center for Neutron Research

National Institute of Standards and Technology

100 Bureau Drive, Bldg 235

Gaithersburg, MD 20899-6102

ABSTRACT

It is common practice to assume that Bragg scattering peaks have Gaussian shape. The Gaussian shape function is used to perform instrumental smearing corrections. Using Monte Carlo ray tracing simulation, the resolution of a realistic small-angle neutron scattering (SANS) instrument is generated reliably. Including a single crystal sample with large d-spacing, Bragg peaks are produced. Their widths are a reliable measure of the standard deviation of the resolution function. Results show that Bragg peaks are not Gaussian for broad neutron wavelength spreads; they are not even symmetric. Moreover, it was found that the universally accepted instrumental smearing procedure used to analyze SANS data yields results that are in error by at least 10 %.

Keywords: Ray tracing simulation, McStas, Small-angle neutron scattering, Bragg peaks, resolution function, Gaussian function

INTRODUCTION

The small-angle neutron scattering (SANS) technique has proven valuable for the characterization of structures in the nanometer size scale. This technique produces over 500 research publications per year. Neutron scattering is limited by modest neutron fluxes due to the inherent nature of neutron sources. In order to enhance flux-on-sample, SANS instruments use velocity selectors (for monochromation) with wide wavelength spreads. This broadens instrumental resolution. SANS data are usually analyzed by fitting to appropriate models that are smeared in order to account for non-negligible resolution effects.

Precise investigations of the SANS resolution function proceed in one of two ways, either experimentally or through simulation.  Measuring scattering from samples that are characterized by Bragg peaks would cover the experimental approach since the width of Bragg peaks is dictated by the instrumental resolution along with the mosaic spread of the sample. Unfortunately, it is hard to find “single-crystal diffraction” samples with nanometer scale d-spacings. Moreover, the mosaic spread is oftentimes non-negligible and not known. This leaves us with the second approach of simulating the resolution function.

Using Monte Carlo ray-tracing simulation, the SANS resolution is critically analyzed here. Using typical instrument configurations, the width of the instrumental resolution is obtained by measuring the standard deviation of the main neutron beam on the area detector. In order to simulate the resolution at finite scattering angle (or finite scattering variable Q), a single-crystal diffraction sample is used to simulate Bragg peaks. The position and standard deviation of these Bragg peaks determine the resolution function under “controlled” sample mosaic spread conditions.

The Q=0 and finite-Q resolution functions are analyzed and compared to analytical predictions.

ANALYTICAL EQUATIONS

The analytical equations used to predict the SANS resolution are reviewed briefly here. These are universally used for all SANS resolution corrections.

SANS instruments contain four major steps: monochromation using velocity selectors, collimation using source and sample apertures, scattering from samples with nanometer size structures and detection using position-sensitive area detectors. In order to resolve nanostructures, long flight paths (typically tens of meters) are often used. The scattering variable is defined in terms of the scattering angle  and neutron wavelength as .

The standard deviation of the resolution function has two main contributions: one from the geometry of the beam and the other from the neutron wavelength spread. Neglecting gravity effects, the horizontal variance is expressed in the radial and perpendicular (i.e., tangential) directions as [1,2]:

(1)

.

The standard instrument parameters are defined as: L1 and L2 are the source-to-sample and sample-to-detector distances, r1 and r2 are the source and sample apertures radii, x3 and y3 are the detector cell dimensions,  is the neutron wavelength and is the FWHM of the triangular wavelength distribution. This variance was calculated as the second moment averaged over source and sample apertures and detector cell as well as over the triangular wavelength distribution.

MONTE CARLO SIMULATIONS

Using the McStas package [3], Monte Carlo ray tracing simulations were performed to simulate a realistic SANS instrument with L1 = 10 m, L2 = 10 m, r1 = 1 cm, r1 = 0.5 cm, and x3 = 0.55 mm (for high spatial resolution).

The direct neutron beam measured by the area detector corresponds to setting Qx = 0 in Eq. (1). Note that the Q-variance and the spatial variance (in the detector plane) are related by. Here, the simulated spatial variance is obtained using one of two methods; either by fitting the simulated data for the direct beam to a 2D Gaussian function or by numerically calculating the second moment of the peaked distribution as  where  is the average over the direct beam spot. These second moments are calculated numerically by summing over the beam spot with detector counts as the weighing factor using the OriginPro software package. Between 107 and 108 neutrons are used for each simulation.

The first result is that the calculated second moments agree well with the analytical results (obtained from Eq (1) with Qx=0), but the fits to the 2D Gaussian results are systematically some 10 % higher than the analytical results. For the instrument configuration considered, the 2D Gaussian result gives for the spatial standard deviation , the second moment method gives and the analytical result based on Eq. 1 gives . This is for the resolution at Q=0.

The direct beam central spot is represented in Figure 1. One can notice that horizontal and vertical cuts do not follow a Gaussian shape. They are closer to a triangular shape. Note that in general, this shape is trapezoidal. In the converging geometry (with L1 = L2 and r1 = 2 r2) considered here they become triangular. This result is solid and was checked for many other conditions.

Figure 1: Expanded view of the direct beam central spot with  = 10 Å,  = 1 Å and no sample.

In order to simulate the finite-Q resolution, the “single-crystal” sample module is used in the McStas simulation. A simple cubic crystalline structure is chosen by setting the (h,k,l) reflections along with the unit cell dimensions (a,b,c) as well as the width of the d-spacing distribution  and the angular “mosaic” spread . The Bragg law relates the d-spacing for a specific reflection dhkl, the scattering angle  and the neutron wavelength  as: . The d-spacings are expressed as: . Considering a simple structure with a = b = c =480 Å (large unit cell), a neutron wavelength = 4 Å, and four reflections corresponding to (h=±1,k=±1,l=0), a series of Bragg peaks are obtained on the simulated area detector. The purpose of this exercise is not to simulate any specific single-crystal structure, but rather to produce well-resolved Bragg peaks. When the single-crystal is perfectly aligned along the neutron beam, only the diffraction peaks are observed with no direct beam signal on the detector. Slightly misaligning the crystal allows the direct beam spot to appear as well. Figure 2 shows a typical case with  = 4 Å,  = 1 Å, a narrow  and a mosaic spread of  = 10 arcmin in which both the main (direct beam) central spot and various Bragg peak can be observed. Note that d-spacing spread and the mosaic spread of a crystal affect the beam divergence. For instance, small mosaic spreads tend to focus the beam while broad ones tend to broaden the beam divergence. The direct beam peak is observed by rocking the crystal while making sure that they Bragg peaks remain in the detector plane. The rightmost Bragg peak (along the horizontal x- axis) is of interest here in order to avoid any gravity effect.

Figure 2: Simulated SANS data with  = 4 Å,  = 1 Å. The central (direct beam) peak and a multitude of Bragg peaks can be resolved at the same time by slightly rocking the crystal.

Expanding the rightmost Bragg peak (in the horizontal direction) one sees in Figure 3 that this finite-Q peak also does not follow a Gaussian shape. Its shape is asymmetric and more stretched towards the high-Q values for the radial (horizontal) direction. The perpendicular (vertical) direction is symmetric.

Figure 3: Expanded horizontal (rightmost in Figure 2) Bragg peak for  = 4 Å,  = 1 Å.

Figure 4 shows a direct comparison of the horizontal cut of the same Bragg peak with the Gaussian shape  in the horizontal direction. There are noticeable deviations. Moreover, the third moment (curtosis) of the simulated data is not zero as it should be for a Gaussian distribution.

Figure 4: Comparison of the horizontal cut across the Bragg peak with the fitted 2D Gaussian shape. The Bragg peak was obtained for  = 4 Å,  = 1 Å.

The standard deviation of the simulated Bragg peak along the horizontal axis is calculated numerically using the 2D Gaussian fit and using the second moment method; then these values are compared to the analytical estimate using Eq. (1) as shown in Figure 5. The Bragg peak position (determining the scattering variable) is varied by changing the neutron wavelength in order to keep the single crystal aligned. Each time the neutron wavelength was changed, the unit cell dimensions (a,b,c) were adjusted proportionately (but keeping a=b=c) in order to keep the Bragg spots in the detector  plane. The spatial standard deviation (units of mm) is used for the ordinate axis. Here also, one can see that the fit to a 2D Gaussian function (red curve in Fig. 5) is systematically between 10 % and 20 % higher than the second moment method (green curve) depending on the Q value. Moreover, the analytical approach (blue curve) agrees well with the second moment method. The d-spacing spread of and the mosaic spread of 10 arcmin used here are reasonable since they reproduce the predicted value of the spatial standard deviation rather well at low-Q. This is the flat (Q-independent) part where the geometric contribution (term in square brackets in Eq. 1) dominates. At higher Q, the wavelength spread becomes important.

Figure 5: Variation of the simulated Bragg peak spatial standard deviation when using the 2D Gaussian fit method and the second moment method and comparison to the analytical predictions for finite-Q. The wavelength is varied, but keeping  = 1 Å. Smooth lines have been added as guide to the eye. Statistical error bars correspond to one standard deviation.

RESULTS AND DISCUSSION

Small-angle neutron scattering is a popular analytical method for the characterization of soft matter including polymers, complex fluids and bio-macromolecules. Resolution corrections are performed assuming a Gaussian shape function with a standard deviation calculated using a universally accepted analytical formalism (Eq. 1). Using Monte Carlo ray tracing simulation of the instrumental resolution, Bragg peaks representing structures with large d-spacings were generated. The shape of Bragg peaks was found to be different from the Gaussian function universally assumed for the resolution function especially for long wavelength spreads (). Bragg peaks tend to be elongated in the radial direction and not symmetric about the peak position. The widths of these Bragg peaks were estimated either using 2D fits to a Gaussian function or directly calculated numerically as second moments. The standard deviation obtained from the 2D Gaussian fits was found to be off by 10 % to 20 % from the more reliable second moment method. Eq. 1 is based on the second moment. This systematic error hampers all published SANS results.

It is common practice to correct for instrumental resolution effects by smearing the model before performing nonlinear least-squares fits to the data. This smearing step consists in modeling the resolution function as a Gaussian function with a standard deviation which is calculated analytically using Eq. (1). In the light of the systematic errors discussed here, it would be wise to reconsider this method. The resolution shape function can be generated rather precisely using Monte Carlo simulation as shown here. Instead of using a generic Gaussian shape, this simulated resolution shape function could be used to perform the resolution integral numerically. That may slow down model fitting but would be more precise.

This scientific contribution is fitting since it falls close in time to the 100-year anniversary of the original Bragg publication [4].

DISCLAIMER/ACKNOWLEDGMENTS

The identification of commercial products does not imply endorsement by the National Institute of Standards and Technology nor does it imply that these are the best for the purpose. This work is based upon activities supported in part by the National Science Foundation under Agreement No. DMR-0944772. Help from Sylvain Desert and Emmanuel Farhi with the MSCTAS package and from Hideo Fujii with the OriginPro software is greatly appreciated.

REFERENCES

[1] Mildner, D.F.R. & Carpenter, J.M. (1984). J. Appl. Cryst. 17, 249-256.

[2] Hammouda, B. and Mildner, D.F.R. (2007). J. Appl. Cryst. 40, 250-259.

[3] Willendrup, P., Farhi, E. & Lehmann, K. (2004). “McStas 1.7- a New Version of the Flexible Monte Carlo Neutron Scattering Package”, Physica B - Condensed Matter 350, E735-E737. The McStas package can be downloaded from http://www.mcstas.org.

[4] Bragg, W.H. (1912). Nature (London) 90, 219