National Institute of Standards & Technology

Certificate of Analysis

##### Standard Reference Material( 3183

## Fluoride Anion Standard Solution

Lot No. 050718

This Standard Reference Material (SRM) is intended as a primary standard for the quantitative determination of fluoride using anion ion chromatography (IC) or other methods.  One unit of SRM 3183 consists of 50 mL of solution in a high density polyethylene bottle sealed in an aluminized bag.  The solution is prepared gravimetrically to contain a known mass fraction of fluoride dissolved in filtered (0.22 m) water having a minimum resistivity of 18  cm.

Certified Value of Fluoride:     1.001 mg/g  (  0.006 mg/g

The certified value is based on (1) gravimetric preparation using high purity sodium fluoride and (2) anion IC using two independently prepared primary standards.  No correction has been applied for transpiration that will occur after the SRM bottle unit has been removed from the sealed bag.  See “Instructions for Use” for more information regarding transpiration.

The uncertainty in the certified value is calculated as

U  =  kuc

where k = 3.18 is the coverage factor for a 95 % confidence interval.  The quantity uc is the combined standard uncertainty calculated according to the ISO and NIST Guides [1].  The value of uc is intended to represent, at the level of one standard deviation, the combined effect of uncertainty components associated with the gravimetric preparation, the IC determination, and method bias [2].

Expiration of Certification:  The certification of SRM 3183 Lot No. 050718 is valid, within the measurement uncertainty specified, until 14 November 2010, provided the SRM is handled in accordance with instructions given in this certificate (see “Instructions for Use”).  This certification is nullified if the SRM is damaged, contaminated, or modified.

Coordination of the technical measurements leading to the certification of SRM 3183 was provided by M.R. Winchester of the NIST Analytical Chemistry Division.

This SRM was prepared gravimetrically and analyzed using IC, and primary standards for IC calibration were prepared and validated, by T.A. Butler and M.S. Rearick of the NIST Analytical Chemistry Division.

Statistical consultation was provided by S.D. Leigh of the NIST Statistical Engineering Division.

The support aspects involved in the issuance of this SRM were coordinated through the NIST Measurement Services Division.

Stephen A. Wise, Chief

Analytical Chemistry Division

Gaithersburg, MD 20899	Robert L. Watters, Jr., Chief

Certificate Issue Date:  18 July 2006	Measurement Services Division

Maintenance of Certification:  NIST will monitor representative solutions from this SRM lot over the period of its certification.  If substantive changes occur that affect the certification before the expiration of certification, NIST will notify the purchaser.  Registration (see attached sheet) will facilitate notification.

TRACEABILITY

Calibration of analytical instruments or procedures for the determination of fluoride should be performed using standards whose values are traceable to the certified value of fluoride in this SRM.  Traceability must be established through an unbroken chain of comparisons, each having stated uncertainties [3].  Comparisons are based on appropriate physical or chemical measurements.  These may include chromatographic or other methods of analysis.  Gravimetric or volumetric dilution is also a method of comparison, where the mass or volume of a solution before and after dilution is measured.  The uncertainties assigned to the traceable values of such standards must include the uncertainty of the certified value of fluoride in this SRM, appropriately combined with the uncertainties of all comparison measurements.

# INSTRUCTIONS FOR USE

This SRM can be used to prepare working standard solutions with fluoride mass fractions in the range 10 mg/kg to 100 mg/kg, from which more dilute standards can be prepared.  The user should establish internal laboratory procedures that specify a maximum shelf-life for a working standard solution.  Two procedures for the preparation of working standard solutions follow.

Preparation of Working Standard Solutions by Mass:  Each working standard solution should be prepared by transferring an aliquot of the SRM into an empty, dry, preweighed, polyethylene bottle and then reweighing the bottle.  Water of appropriate purity must be added by mass to bring the solution to the desired dilution.  The dilution need not be exact since the mass of the empty bottle, mass of the bottle plus SRM aliquot, and the final diluted mass of the solution will permit calculation of the exact mass fraction (i.e., mass of fluoride per mass of solution) of the working standard solution.  Dilutions prepared gravimetrically as described will need no correction for temperature and no further correction for true mass fraction in vacuum.  Volumetric dilutions are NOT recommended due to uncertainties in volume calibrations and variations in density.  However, for user convenience, a procedure for volumetric preparation that will minimize the major sources of error is given below.

Preparation of Working Standard Solutions by Volume:  Each working standard solution should be prepared by transferring an aliquot of the SRM into an empty, dry, polyethylene bottle and then weighing the bottle.  The solution must now be transferred to a Class A volumetric flask, and the polyethylene bottle must be reweighed to determine the exact mass of SRM solution transferred.  The solution in the flask is then diluted to 99 % + volume using water of appropriate purity, mixed thoroughly, and the remaining few drops needed to dilute to exact volume carefully added.  The concentration (in mg/mL) of the resulting working standard solution can then be calculated by multiplying the mass (in g) of the SRM solution amount by the SRM certified value (in mg/g) and dividing the numerical product by the calibrated volume (in mL) of the flask used for dilution.  Thus, no correction for density is needed.  Although the concentration of the resulting working standard solution may be an uneven fraction of the original SRM concentration, it will be known as accurately as a volumetric dilution permits.

Transpiration:  While stored in the sealed aluminized bag, transpiration of this SRM is negligible.  After the SRM has been removed from the aluminized bag, transpiration will occur at a solution mass loss rate of 
approximately 0.2 % relative per year, resulting in a gradual increase of the mass fraction.  It is the responsibility of the user to account for this effect.  The recommended way to reduce the effect of transpiration is to deliver all of the SRM as aliquots weighed into appropriate vessels as soon as the SRM is removed from the aluminized bag.  The aliquots may be stored and can be diluted to known mass or volume at a later date.  Storage of a partially used SRM bottle is NOT recommended; however, if such storage is necessary, the cap should be tightly sealed and the SRM bottle kept in an airtight container to slow the rate of transpiration.  When the bottle is weighed both before and after being placed in storage, the mass difference observed will be a measure of transpiration mass loss.  The user should set a maximum shelf-life for a partially used SRM bottle commensurate with accuracy requirements.

REFERENCES

[1]	ISO; Guide to the Expression of Uncertainty in Measurement; ISBN 92-67-10188-9, 1st ed., International Organization for Standardization:  Geneva, Switzerland (1993); see also Taylor, B.N.; Kuyatt, C.E.; Guidelines for Evaluating and Expressing the Uncertainty of NIST Measurement Results, NIST Technical Note 1297, U.S. Government Printing Office: Washington, DC (1994); available at http://physics.nist.gov/Pubs/.

[2]	Levenson, M.S.; Banks, D.L.; Eberhardt, K.R.; Gill, L.M.; Guthrie, W.F.; Liu, H.K.; Vangel, M.G.; Yen, J.H.; Zhang, N.F.; An Approach to Combining Results From Multiple Methods Motivated by the ISO GUM; J. Res. Natl. Inst. Stand. Technol., Vol. 105; p. 571 (2000).

[3]	International Vocabulary of Basic and General Terms in Metrology; ISBN 92-67-01075-1, 2nd ed., ISO, Geneva, Switzerland (1993).

Users of this SRM should ensure that the certificate in their possession is current.  This can be accomplished by contacting the SRM Program at:  telephone (301) 975-6776; fax (301) 926-4751; e-mail srminfo@nist.gov; or via the Internetat  http://www.nist.gov/srm.